## Slide 1: NeuroDOT
Tutorial:
Data Quality, Pre-processing, and Reconstruction

## Slide 2
NeuroDOT Set-Up - Checklist
Before starting any tutorial, please make sure you’ve done the following
Download the appropriate versions of matlab, neuroDOT, and NIRFAST
  Matlab: 2020b
  NeuroDOT toolbox from NITRC
    https://www.nitrc.org/projects/neurodot/
  NIRFASTer-master
    https://github.com/nirfaster/NIRFASTer
Set your path so that all required folders and subfolders are included
  Folders w/subfolders to include
    NeuroDOT toolbox
    NIRFASTer-master
    Your personal directory
  Folders to include
    Output directory

## Slide 3
NeuroDOT Set-Up – Output Directory
Once you’ve set your path, the next thing you should do is set your output directory.
  The output directory must be located within your path
  If it isn’t already in your path, you must add it to your path
You can set your own path to an output directory in the line shown above
An output directory specifies the location where files will be saved
  MATLAB, by default, will save files to the current directory
The following line of code will change your current directory to the output directory that you just specified 
Now you can save outputs directly to your output directory
Example output directory

## Slide 4: Diffuse Optical Tomography
In diffuse optical tomography (DOT), arrays of optical sources and detectors are used to perform functional neuroimaging experiments on task-based or resting state modes of human brain function.
Analyses are broken into several pipelines:
  B: modeling of the tissue shape, optical property distribution, and source/detector locations
  C: modeling of the light emission, diffusion, and detection through the head
  D: preprocessing of the raw source-detector measurements
  E: reconstruction and spectroscopy of the preprocessed data and light model into a functional neuroimaging volume
  F: post-processing analysis of these results.

## Slide 5: Pre-processing and Image Reconstruction
This tutorial will follow the NeuroDOT_Visualization _Script_Tutorial.m Matlab file which can be found in the Documentation/Scripts folder.
This version of the Visualization Overview tutorial utilizes the CCW1 data sample which can be found in the Data/ folder of the NeuroDOT toolbox
The tutorial will:
  Generate raw data quality assessment figures
  Process and visualize source-detector measurements
  Perform image reconstruction using a light model sensitivity matrix A
  Visualize the reconstructed data
We will use data from a rotating-wedge retinotopy experiment

## Slide 6: Pre-processing and Image Reconstruction
To fully processes the Sample Datasets contained within the toolbox, you will need to generate the Sensitivity matrices for the 24x28 and 96x92 (source x detector) arrays.
If these Sensitivity matrices have not yet been generated, you can make them using the files and scripts included in the toolbox.
To do so, please see the NeuroDOT_Tutorial_Generating_a_Light_Model_with_AlignMe.pptx and its associated script for the 24x28 array and the …Pad_Adult_96x92_Example.pptx and its associated script for the 96x92 array.
After generating, be sure to add the name of the relevant Sensitivity matrix file to the %% Load Measurement data section of the NeuroDOT_Full_Processing_Script_Tutorial.m script.

## Slide 7
Running the script – Author’s Note
How the author runs the script
  Highlight and execute each chunk of code
    Ex: highlight lines 11 and 12  right click  hit “Evaluate Selection”
Alternate way of running the script
  For each section (beige) hit ctrl+enter
  This isn’t how the author does it, but it will still work
Note: DO NOT HIT THE GREEN “RUN” BUTTON to run this script
Example
section

## Slide 8: NeuroDOT_Data_Sample_CCW1.mat
The CCW1 data set is one of those acquired for the Eggebrecht et al., Neuroimage, 2012 paper.
Briefly, the participant gazed upon a fixation cross while a reversing checkerboard progressed in the counter-clockwise (CCW) direction.
In the figures below, synchronization points labelled with a red line denote resting epochs; those with a green line denote the start of the CCW block wherein the flickering wedge is located at 6:00.
Please see the paper for more information.
Matlab code segments that can be copied and pasted into matlab to be run are in Courier BOLD font.
dataset='CCW1';
load(['NeuroDOT_Data_Sample_',dataset,'.mat']); % data, info, flags
% Set parameters for A and block length for quick processing examples
switch dataset
case {'CCW1','CCW2','CW1','OUT'}
A_fn='A_AdultV24x28.mat'; % Sensitivity Matrix
dt=36; % Block length
tp=16; % Example (block averaged) time point
case {'IN1'}
A_fn='A_AdultV24x28.mat'; % Sensitivity Matrix
dt=36; % Block length
tp=32; % Example (block averaged) time point
case {'HW1','HW2','RW1','GV1','HW3_Noisy'}
A_fn='A_Adult_96x92_on_Example_Mesh_test.mat'; % Sensitivity Matrix
dt=30; % Block length
tp=16; % Example (block averaged) time point
end

## Slide 9: Flow Chart for Pre-processing and Image Reconstruction
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Spatial normalization
Light
Model

## Slide 10: Raw Data Quality Assessment: Time Traces
The left panels show data at 750 nm and the right panels show data at 850 nm.
The top panels show all measurements and the bottom panels measurements with standard deviation below a set threshold
The y-axis on all is log scale to highlight the range of light levels for the set of source-detector distances.
Vertical lines on the bottom row of plots signify time points where something happened in the stimulus paradigm.
Plot_RawData_Time_Traces_Overview(data,info);
DOT Acquisition

## Slide 11: Raw Data Quality Assessment: Cap-relevant views
DOT Acquisition
Top row shows average light levels for short-distance (left) and medium-distance (right) measurements. Good coupling across the entire cap results in mostly white and yellow in these figures.
Middle row shows signal-to-noise ratio (SNR) for cardiac pulse band of frequencies.
  Ideally, the maximum SNR will be greater than 20 for these two plots
Bottom edge plot shows measurements that satisfy the criteria. Optodes that have multiple ‘bad’ measurements are identified with large circles.
Plot_RawData_Cap_DQC(data,info);

## Slide 12: Raw Data Quality Assessment: The signal and the noise
DOT Acquisition
(Upper left) A zoomed-in view of several signals shows the presence of cardiac pulse at approximately 1 Hz.
(Lower left) Spectral measures averaged over several time traces show strong peaks at ~1Hz (cardiac pulse) and approximately 1/26 Hz (the stimulus frequency).
(Upper right) Light level (log) vs. source-detector distance shows expected falloff, with cross-talk at distances > 45 mm.
(Lower right) Histogram of signal % standard deviation shows several measurements below ‘good measurement’ threshold of 7.5%.
Plot_RawData_Metrics_I_DQC(data,info)

## Slide 13: Raw Data Quality Assessment: The signal and the noise
DOT Acquisition
A zoomed-in view of several signals shows the presence of cardiac pulse at approximately 1 Hz.
Approximately 1 second passes between the peaks of each pulse
Plot_RawData_Metrics_I_DQC(data,info)

## Slide 14: Raw Data Quality Assessment: The signal and the noise
DOT Acquisition
Spectral measures averaged over several time traces show strong peaks at frequencies associated with important phenomenon
Peak 1 @ 1Hz = cardiac pulse
  Heart rate
Peak 2 @ 1/26 HZ = stimulus frequency
  This should be the largest peak visible, and is indicative of low motion/good data quality
Plot_RawData_Metrics_I_DQC(data,info)
Peak 1
Peak 2

## Slide 15: Raw Data Quality Assessment: The signal and the noise
DOT Acquisition
Light level (log) vs. source-detector distance shows expected falloff, with cross-talk at distances > 45 mm.
For good data
  Bands are approximately 2 orders of magnitude in length
  There’s a log-linear trend to the falloff plot
    Bars will follow a decreasing linear line, like the arrow shown in green. Ideally, the middle of each bar will be along the line
    The plot on the right doesn’t have a great log-linear falloff, but this is a good example of a mediocre falloff plot
Plot_RawData_Metrics_I_DQC(data,info)

## Slide 16: Raw Data Quality Assessment: The signal and the noise
DOT Acquisition
Histogram of signal % standard deviation shows several measurements below ‘good measurement’ threshold of 7.5%.
  Red line = threshold
Good data will have all, or almost all, measurements to the LEFT of the red line.
Poor data will have many measurements to the RIGHT of the red line
Plot_RawData_Metrics_I_DQC(data,info)

## Slide 17: Pre-processing
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Pre-processing generally follows the pipeline on the right, code shown below:
  lmdata = logmean(data); % Logmean Light Levels
  info = FindGoodMeas(lmdata, info, 0.075); % Detect Noisy Channels
  lmdata = detrend_tts(lmdata); % Detrend Data
  lmdata = highpass(lmdata, .02, info.system.framerate); % High Pass Filter
  lmdata = lowpass(lmdata, 1, info.system.framerate); % Low Pass Filter 1
  hem = gethem(lmdata, info); % Superficial Signal Regression
  [lmdata, ~] = regcorr(lmdata, info, hem);
  lmdata = lowpass(lmdata, 0.5, info.system.framerate); % Low Pass Filter 2
  [lmdata, info] = resample_tts(lmdata, info, 1, 1e-5); % 1 Hz Resampling
Pre-processing can be broken down into two chunks, pre-filtering, and filtering. Both of these stages will be visualized in the following slides
  For this tutorial, run lines 44-57 to generate both pre-filtered and filtered data, so you can visualize the effects of filtering side-by-side.
Subsequent plots and images show differential measurements, computed as the logarithm of the ratio of each signal and its temporal average
For a more in-depth exploration of pre-processing parameter settings, please see the Tutorial on Pre-processing.

## Slide 18: Pre-processing: pre-filtering
Data used: log-meaned, unfiltered data
  Data has gone through the first 3 steps of pre-processing, but has yet to undergo filtering
Fig 1
  Top: raw data plot
  Middle: image plot, data is scaled to use the full ‘jet’ colormap
  Bottom: Power spectrum
    Before filtering, pulse peak is visible, but it’s hard to make out where the stimulus peak is
    Stim peak is in between and
Fig 2
  Top: GVTD plot
    Measures motion
    Ideally, this plot would have one thin trace with a couple peaks
    More peaks = more motion
  Bottom: Gray, image plot, similar to middle plot of fig 1
    Organized by radius between source and detectors for each measurement
    Red lines = start/stop of stimulus
    Green lines = when stimulus is presented to participant
Fig 1
Fig 2
Pulse Peak

## Slide 19: Pre-processing: post-filtering
Data used: log-meaned, filtered, and resampled data
  Data has gone through all pre-processing steps
Fig 1
  Top: raw data plot
  Middle: image plot, data is scaled to use the full ‘jet’ colormap
  Bottom: Power spectrum
    After filtering, stimulus peak is very evident
Fig 2
  Top: GVTD plot
    Measures motion
    Ideally, this plot would have one thin trace with a couple peaks
    More peaks = more motion
  Bottom: Gray, image plot, similar to middle plot of fig 1
    Organized by radius between source and detectors for each measurement
    Red lines = start/stop of stimulus
    Green lines = when stimulus is presented to participant
Fig 1
Fig 2
Stim Peak

## Slide 20: View effects of Pre-processing side-by-side
(Top) raw data plot
  Notice that after filtering, the length of time shrinks. This is because we filtered out both low, and high frequency noise
(Middle) image plot of scaled data
  Notice that both the time (x axis) and range of the data (colorscale) are lower after filtering
  This is because noise is removed during filtering
(Bottom) Fourier Transformed Data
  Generates an average of the top plot, but on a log-scaled x axis
  Notice that after filtering the data, signal below and above is cropped, allowing for a better focus on brain activity
After Filtering
Before Filtering

## Slide 21: View effects of Pre-processing side-by-side Gray plots
nlrGrayPlots_180818(lmdata,info); % Gray Plot with synch points
After filtering
Before filtering
Peaks in the top plot and vertical lines in the bottom plot denote times where there is high noise
  Before filtering plot has a lot of these
After filtering, high and low frequency noise has been removed, so there are only a few of those vertical lines
  The noise that remains is due to participant motion

## Slide 22: View effects of Pre-processing Wavelength comparison
Plot_TimeTrace_With_PowerSpectrum(lmdata,info); (right)
GrayPlots_Rsd_by_Wavelength(lmdata,info); (below)
Using the same plots, we can take a look at how pre-processing effects both wavelengths of light used.

## Slide 23: View effects of Pre-processing Block averaging
%% Block Averaging the measurement data and view
badata = BlockAverage(lmdata, info.paradigm.synchpts(info.paradigm.Pulse_2), dt);
badata=bsxfun(@minus,badata,mean(badata,2));
figure('Position',[100 100 550 780])
subplot(2,1,1);
plot(badata(keep,:)');
set(gca,'XLimSpec','tight'), xlabel('Time (samples)'), …
ylabel('log(\phi/\phi_0)')
m=max(max(abs(badata(keep,:))));
subplot(2,1,2);
imagesc(badata(keep,:),[-1,1].*m);
colorbar('Location','northoutside');
xlabel('Time (samples)');
ylabel('Measurement #')
After pre-processing, data gets block averaged
Left image: pre-processed
Right image: pre-processed + block averaged
  Top plot: BA data traces
  Bottom plot: image plot of BA data
  Note: this is a cropped version of figure 1 from slide 19

## Slide 24: Reconstruction & Spectroscopy
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Light
Model
Absorption image volumes are reconstructed from the measurements based on a regularized inversion of the A matrix (for details, see, e.g., Eggebrecht et al., Neuroimage, 2012)
  %% Reconstruction Pipeline
  if ~exist('A', 'var') % In case running by hand or re-running script
  A=load([A_fn],'info','A');
  if length(size(A.A))>2 % A data structure [wl X meas X vox]-->[meas X vox]
  [Nwl,Nmeas,Nvox]=size(A.A);
  A.A=reshape(permute(A.A,[2,1,3]),Nwl*Nmeas,Nvox);
  end
  end
  Nvox=size(A.A,2);
  Nt=size(lmdata,2);
  cortex_mu_a=zeros(Nvox,Nt,2);
  for j = 1:2
  keep = (info.pairs.WL == j) & (info.pairs.r2d <= 40) & info.MEAS.GI;
  disp('> Inverting A')
  iA = Tikhonov_invert_Amat(A.A(keep, :), 0.01, 0.1); % Invert A-Matrix
  disp('> Smoothing iA')
  iA = smooth_Amat(iA, A.info.tissue.dim, 3); % Smooth Inverted A-Matrix
  cortex_mu_a(:, :, j) = reconstruct_img(lmdata(keep, :), iA);% Reconstruct Image Volume
  end
  Image volumes store estimated absorption values as voxels by time by wavelength.
  Hemoglobin images are computed from the wavelength-dependent absorption images
  %% Spectroscopy, convert to Hb
  if ~exist('E', 'var'),load('E.mat'),end
  cortex_Hb = spectroscopy_img(cortex_mu_a, E);
  cortex_HbO = cortex_Hb(:, :, 1);
  cortex_HbR = cortex_Hb(:, :, 2);
  cortex_HbT = cortex_HbO + cortex_HbR;

## Slide 25: Spatial normalization
DOT Acquisition
Log-mean
Detect noisy channels
Band-pass filter
Superficial signal regression
Resample
Low-pass filter
Reconstruction
Spectroscopy
Spatial normalization
Light
Model
The cortex_HbO data is in an array that is voxels by time where the voxels are within a subspace described by info.tissue.dim which itself is a subspace of the original anatomical volume used to generate the light model.
To properly contextualize the reconstructed data, it is helpful to view on an appropriate anatomical background.
The following lines will load in the anatomical data used for the light model and spatially register it to the reconstructed optical data.
[MNI,infoB]=LoadVolumetricData('Segmented_MNI152nl_on_MNI111_nifti',[],‘nii');
MNI_dim = affine3d_img(MNI,infoB,A.info.tissue.dim,eye(4),'nearest');
We are using nearest-neighbor interpolation here because we are transforming a mask, so we do not want the values to change.
The next lines will block average the HbO data and put it in the full voxel space to get it ready for visualization
% Block Average Data
badata_HbO = BlockAverage(cortex_HbO, info.paradigm.synchpts(info.paradigm.Pulse_2), dt);
badata_HbO=bsxfun(@minus,badata_HbO,badata_HbO(:,1));
badata_HbOvol = Good_Vox2vol(badata_HbO,A.info.tissue.dim);
tp_Eg=squeeze(badata_HbOvol(:,:,:,tp));

## Slide 26: Visualization of Reconstructed Volumetric Images Using PlotSlices
To get a sense of how to visualize the volumetric info try first with the anatomical volume, use PlotSlices. Note that PlotSlices defaults to an interactive mode. Click in any view plane to browse the reconstructed volume.
  When a slice is clicked, all of the slices update to the new coordinates based on the location clicked
    Try to match your X, Y, Z positions to the images on the right 
  Right click to reset to the original set of slices
  Either middle mouse button (scroll wheel), or Q can be pressed to exit PlotSlices
  Code will not continue until PlotSlices has been quit
(Top) Anatomy only
(Middle) Anatomy + volumetric data
(Bottom) Anatomy + volumetric data + functional data
PlotSlices(MNI_dim)
PlotSlices(MNI_dim,A.info.tissue.dim)
PlotSlices(MNI_dim,A.info.tissue.dim,[], tp_Eg)

## Slide 27: Exploring PlotSlices
To get a sense of how to visualize the volumetric info try first with the anatomical volume, use PlotSlices. Note that PlotSlices defaults to an interactive mode. Click in any view plane to browse the reconstructed volume.
When a slice is clicked, all of the slices update to the new coordinates based on the location clicked
  Try to match your X, Y, Z positions to the images on the right 
  Once you’ve matched your PlotSlices to the first image on this slide, don’t quit PlotSlices
  Keep going and match the next image
Right click on the plot to reset to the original set of slices
Quitting PlotSlices
  Either middle mouse button (scroll wheel), or Q can be pressed to exit PlotSlices
  The black crosshairs will disappear from the plots once you’ve done this, then you’ll be able to close the figure window
  Code will not continue until PlotSlices has been quit
1.
2.
3.
4.

## Slide 28: Exploring PlotSlices
If no parameter structure is given as an input, PlotSlices will generate the below visualization parameters internally
PlotSlices with no parameters input
Full list of parameters that can be used with PlotSlices
  PD: positive definite, set to 1 or 0
  CMAP: colormap
    P/N: can set two different colormaps, one for positive, one for negative
  Th: threshold, contains 2 subfields for pos/neg
    Params.Th.P and Params.Th.N
  Scale: data is scaled using this parameter
  BG: background color
  fig_size: width and height of figure
  fig_handle: actual figure variable
  CH: crosshair on/off, set to 1 or 0
  Cboff: colorbar off, set to 1 or 0
  Cbmode: colorbar params on, set to 1 or 0
  Cbticks: colorbar tick placements
  Cblabels: colorbar tick labels
  Orientation: orientation of volume, transverse (‘t’) or sagittal (‘s’)
  Slices: starting slice numbers to visualize
  Slices_type: are slices in index or coordinate form?

## Slide 29: Exploring PlotSlices
Lets play around with some parameters and see how they affect the visualizations produced by PlotSlices
No params input
Params1 – Positive Definite On
  Params1.PD = 1
Params2 – Thresholds (+/-) On
  Params2.Th.P = 0.001
  Params2.Th.N = -Params2.Th.P
Params3 – Thresholds set to different values
  Params2.Th.P = 0.001
  Params2.Th.N = -0.004
Params1
Params2
Params3
No Params

## Slide 30: Exploring PlotSlices
Let's play around with some parameters and see how they affect the visualizations produced by PlotSlices
No params input
Params4 – Scale
  Params4.Scale = 2*max(abs(tp_Eg(:)))
  Usually, we set the scale to 0.8*max(data) or less than that
Params5 – Colormap
  Params5.Cmap = ‘cool’
Params6 – Positive and Negative Colormaps
  Params6.Cmap.P = ‘winter’
  Params6.Cmap.N = ‘summer’
Params4
Params5
Params6
No Params

## Slide 31: Exploring PlotSlices
Let's play around with some parameters and see how they affect the visualizations produced by PlotSlices
No params input
Params7 – Colorbar Ticks
  Params7.cbmode = 1
  Params7.cbticks = [0,0.25,0.5,0.75,1]
Params8 – Colorbar Labels
  Params8.cbmode = 1
  Params8.cblabels = {‘Hello’, ‘Goodbye’}
Params 9– Put it all together
  Params9.cbmode = 1
  Params9.cbticks = [0,0.25,0.5,0.75,1]
  Params9.cblabels = {‘1’, ‘2’, ‘3’, ‘4’, ‘5’}
  Params9.Cmap.P = ‘winter’
  Params9.Cmap.N = ‘summer’
  Params9.Scale = 0.8*max(abs(tp_Eg(:)))
  Params9.Th.P = 0.001
  Params9.Th.N = -0.004
Params7
Params8
Params9
No Params

## Slide 32: Visualization of Reconstructed Volumetric Images
Now that you’ve gotten acquainted with PlotSlices, lets take a look at the data. First, with the data by itself, then on top of the underlay and volumetric info. Then, set scaling and thresholding parameters to home in on strong activations and visualize again.
Full code section
% Visualize the data
PlotSlices(tp_Eg,A.info.tissue.dim); %data by itself
PlotSlices(MNI_dim,A.info.tissue.dim,[],tp_Eg); %data with anatomical underlay
% Set parameters to visualize more specific aspects of data
Params.Scale=0.8*max(abs(tp_Eg(:))); %scale wrt/max of data
Params.Th.P=0.5*Params.Scale; %threshold to see strong activations
Params.Th.N=-Params.Th.P; %thresholds go both ways
Params.Cmap='jet';
PlotSlices(MNI_dim,A.info.tissue.dim,Params,tp_Eg);
Individual visualization lines
(Top) Functional data and volumetric info only
(Middle)
(Bottom)
PlotSlices(tp_Eg,A.info.tissue.dim);
PlotSlices(tp_Eg,A.info.tissue.dim, [], tp_Eg);
PlotSlices(tp_Eg,A.info.tissue.dim, Params, tp_Eg);
Match slices to replicate these visualizations

## Slide 33: Visualization with Voxel Time Trace
We can also visualize data with its time trace. Click on an activation site to view the time trace for that slice. Then, click on different points of the time trace to se how the activation pattern changes over the course of the CCW task
Adding the time trace at the selected location further aids interpretation
% Explore the block-averaged data a bit more interactively
Params.Scale=0.8*max(abs(badata_HbOvol(:)));
Params.Th.P=0;
Params.Th.N=-Params.Th.P;
PlotSlicesTimeTrace(MNI_dim,A.info.tissue.dim,Params,badata_HbOvol,info)
  In this data set, regions within the visual cortex show varying time to peak response, corresponding to the varied position of the rotating checkerboard wedge.
  These data are block-averaged, which reduces variance in the reconstructed signals. Visualizing the un-averaged data (next) can help in interpretation.
% Explore the not-block-averaged data a bit more interactively
HbOvol = Good_Vox2vol(cortex_HbO,A.info.tissue.dim);
Params.Scale=2e-3;
Params.Th.P=1e-3;
Params.Th.N=-Params.Th.P;
PlotSlicesTimeTrace(MNI_dim,A.info.tissue.dim,Params,HbOvol,info)
Match slices and time traces
to replicate these visualizations

## Slide 34: Visualization of reconstructed data on a surface model of the cortex
We can also visualize functional data on a 3D model of the brain
%% Select Surface visualizations
if ~exist('MNIl', 'var'),load(['MNI164k_big.mat']);end
HbO_atlas = affine3d_img(badata_HbOvol,A.info.tissue.dim,infoB,eye(4));
tp_Eg_atlas=squeeze(HbO_atlas(:,:,:,tp));
pS=Params;
pS.view='post’; % Posterior view
pS.ctx='std'; % Standard pial cortical view
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
Parameters used
Scale = 0.004
Thresholds
  P: 0.001
  N: -0.001
Colormap = jet
View = ‘post’ (posterior)
Inflation (ctx) = ‘std’ (standard)

## Slide 35: Visualization of reconstructed data on a surface model of the cortex – exploring parameters
Let’s play around with some parameters and explore how they affect the resultant visualizations.
View
Dorsal view
Lateral view
Inflation (back to posterior view for this)
Standard
Inflated
Very Inflated
Dorsal
Lateral
pS.view=‘dorsal’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.view=‘lat’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
Standard
Inflated
Very Inflated
pS.ctx=‘std’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.ctx=‘inf’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.ctx=‘vinf’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);

## Slide 36: Visualization of reconstructed data on a surface model of the cortex – exploring parameters
Let’s play around with some parameters and explore how they affect the resultant visualizations.
Thresholds
Both at zero
Set a bit higher
Set even higher
Set very high
pS.Th.P = 0;
pS.Th.N = -pS.Th.P;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.Th.P = 1e-4;
pS.Th.N = -pS.Th.P;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.Th.P = 1e-3;
pS.Th.N = -pS.Th.P;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.Th.P = 1e-2;
pS.Th.N = -pS.Th.P;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
Th = 0
Th = +/- 1e-4
Th = +/- 1e-3
Th = +/- 1e-2
Note:
As the thresholds increase, less data is visualized. When setting thresholds for your own data, we recommend finding thresholds that allow for the best focus on activation

## Slide 37: Visualization of reconstructed data on a surface model of the cortex – exploring parameters
Let’s play around with some parameters and explore how they affect the resultant visualizations.
Scale
Start with 4e-3
Lower to 1e-3
Raise to 9e-3
Lowering the scale will increase saturation of activation sites, while raising it will decrease the saturation
pS.Scale = 4e-3;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.Scale = 1e-3;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS.Scale = 9e-3;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
Scale = 1e-3
Scale = 4e-3
Scale = 9e-3

## Slide 38: Visualization of reconstructed data on a surface model of the cortex – exploring parameters
Let’s play around with some parameters and explore how they affect the resultant visualizations.
Colormaps
Same colormap for positive and negative
Different colormaps for positive and negative
All parameters used in PlotSlices can be used with PlotInterpSurfMesh
  PlotInterpSurfMesh = NeuroDOT function that allows for visualizing functional data on a cortical mesh
pS.Cmap = ‘cool’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
pS = rmfield(pS, 'Cmap');
pS.Cmap.P = ‘winter’;
pS.Cmap.N = ‘summer’;
PlotInterpSurfMesh(tp_Eg_atlas, MNIl,MNIr, infoB, pS);
Single colormap
Dual colormaps

## Slide 39: Visualization of reconstructed data on a surface model of the cortex – exploring parameters
Let’s play around with some parameters and explore how they affect the resultant visualizations.
% Get data at another timepoint
tp_Eg_atlas_2=squeeze(HbO_atlas(:,:,:,tp+9));
% Make masks of both time points
mask_temp_1 = tp_Eg_atlas;
mask1 = zeros(size(mask_temp));
mask1(mask_temp > pS.Th.P) = 1; %set tp1 positive activations to 1
PlotInterpSurfMesh(mask1, MNIl,MNIr, infoB, pS); %visualize tp1
mask_temp_2 = tp_Eg_atlas_2;
mask2 = zeros(size(mask_temp_2));
mask2(mask_temp_2 > pS.Th.P) = 1; %set tp2 positive activations to 1
PlotInterpSurfMesh(mask2, MNIl,MNIr, infoB, pS); %visualize tp2
% Combine masks to plot
mask3 = zeros(size(mask_temp));
mask3(mask1 == 1 & mask2 == 1) = 3;
mask3(mask1 == 1 & mask3 ~= 3) = 1;
mask3(mask2 == 1 & mask3 ~= 3) = 2;
% Set parameters and visualize
pS_TC = pS;
pS_TC.Cmap = [1,0,0; 0,0,1; 1,0,1]; %make a colormap
pS_TC.TC = 1; %True Color On
pS_TC.OL = 1; %Overlap On
PlotInterpSurfMesh(mask3, MNIl,MNIr, infoB, pS_TC);
Mask for tp1
Mask for tp2
Overlap
We can use the TrueColor Parameter, TC to overlap maps of data from different time points, subjects, scans, etc. on the same brain.
Here we take only the positive activations (activations greater than our threshold of 1e-3) from two time points of the CCW1 sample data and create binary masks for those activations and visualize those
Then, we combine those masks into one that contains info on where activations from both time points do and do not overlap.
Finally, we create a custom colormap where:
Red = activations from tp1
Blue = activations from tp2
Magenta = activations from where tp1 and tp2 overlap

## Slide 40: Help Expand NeuroDOT’s Utility!
Please fill out the NeuroDOT registration form to help us better develop the NeuroDOT toolbox and expand its utility to address your fNIRS/DOT data analysis needs:
  https://forms.gle/8QNGnx7ZbKuUHg3bA
Additionally, please provide specific feedback, or ask questions to the development team:
  https://forms.gle/jv6RkX5s784LgQC89
Note: you will not be contacted by the NeuroDOT development team unless you opt-in to receiving communications.

## Slide 41: That’s It (For Now)
Congratulations! You have finished the NeuroDOT Tutorial for the full processing of CCW data acquired with the 24x28 visual pad.
See other Sample Results Appendices for example visualizations of these same processing steps on the other Sample Data included in the toolbox.
Also, see the PreProcessing and Reconstruction tutorials to gain a deeper understanding of the effects of altering the processing parameters.
NeuroDOT Support: neurodot-support@wustl.edu